Approvals

| Title | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Gary Chopcinski |  |  |
| Business Unit Manager/ Director | Lisa Rotelli |  |  |
| Director of New Product Development | James Pham |  |  |
| Director of Product Risk & Quality Engineering | Joe Cervantez |  |  |
| Manager of Regulatory Affairs | Lisa Rotelli |  |  |

Change History

| Revision | Revision Date | Name | Description of Change |
| --- | --- | --- | --- |
| A | May 7, 2019 | Gary Chopcinski | Initial Release |

# Introduction

## Scope

This document defines the requirements for the ASL Proportional Head Array Control, a.k.a ASL110.

## Definitions

### ASL110

“ASL110” refers to the Head Array driver control connects to a system using the generic 9-pin DSUB connection.

### User

The patient becomes the user when he or she controls the power wheelchair through the driver control. By definition, there is only one user of the power wheelchair at any one time.

### Patient

The patient is the individual for whom the power wheelchair is prescribed, seated in and controlling the power wheelchair through the driver control input device.

### Caregiver

The caregiver is anyone who assists patient with transfer into/out of the power wheelchair or assists with the operation of/or positioning the power wheelchair. The caregiver either controls the motion of the power wheelchair manually or using the electronic motion controls.

### Provider

The “Provider” is a person or organization which supplies and maintains the items on behalf of the Patient.

### Power Mobility Device (PMD)

A PMD is a mobility device with direct drive motors to propel the drive wheels, on board batteries, a control module to send electric current to the motors and a control input device (joystick, driver control) to send directional and speed commands to the control module.

### Expandable Electronics

An electronic system capable of accommodating three or more actuators if enabled, accepting digital and other alternative driver controls, and operating other devices through the driver control input in addition to the power wheelchair.

### Powered Seating

Actuator controlled seating that changes the body's position in space.

# References

- Article for Power Mobility Devices - Policy Article - Effective January 2009 (A47122)

- Wheelchair Options/Accessories - Policy Article - Effective January 2010 (A47229)

- RESNA WC-2:2009

- Local Coverage Determination (LCD) for Power Mobility Devices (L27239)

- Local Coverage Determination (LCD) for Wheelchair Options/Accessories (L27223)

# Product Description

The ASL110 Electronic Head Array system has electronics and sensors to allow the ASL110 to control wheelchair features including, but not limited to, driving, power positioning and connectivity. The ASL110 is used by the patient’s head to provide the equivalent to Joystick Demand, therefore, the ASL110 must be mounted to have access to the patient’s head.

The ASL110 will have 3 directions of control/joystick demand; left, right and forward. The “forward” sensor is located directly behind the patient’s head and the left and right sensors are located to the right and left of the patient’s head. The mounting will have appropriate adjustability to accommodate the typical patient’s head dimensions.

The ASL110 will interface to a Display device for calibration and setup of the ASL110.

There will be a programmable timer to adjust the reaction of the ASL110 to the patient’s capabilities.

Power On/Off: The on/off switch is located right on the back of the device accessible by the Caregiver but not necessarily to the Patient.

There is a Jack Socket to plug in a mono switch (Egg Switch) for control of the ASL110 by the patient. This button will provide a short and long press capability to provide two levels of access. The short will execute the active and the long press will advance to the next feature.

This provides the patient with the ability to turn the ASL110 on and off without affecting the controlling system.

# Product Features

## Sensors

There will be 3 proportional sensors and 3 digital sensors. One of each type of sensor shall be mounted in single “pads”.

## Display Device

A hand-held programmer for setup, calibration and diagnostics. This device has a Color Graphic screen, up button, down button and a cord to plug into main part of the ASL110.

## Mounting

Mounting compatible with existing Invacare wheelchairs.

The pads (with sensors) will mounted near the back of the head, one on the left side of the head and one on the right side of the head.

# User Characteristics, Interfaces and Environment

## User Characteristics

### Patient, User

This alternative driver control system is typically used for individuals with Cerebral Palsy or with anyone that has difficulty with motor control, coordination or weakness.  This patient has a certain degree of head control sufficient to operate low/no pressure switches mounted at the left, right and back of the head.

### Caregiver

The Caregiver is a person available during most of the day and night and can be a hired individual, a family member or friend.  This person can be medically educated to the Patient’s needs or not.

### Provider

The provider is an individual, company or institution supplying the equipment to the Patient.  The provider generally provides instructions for use and maintenance support of the equipment.

## User Interfaces

### Patient

The patient uses the ASL110 sensor pads to issue “joystick demand” to the Wheelchair’s Control System.

### Caregiver

The ASL110 must mount to the wheelchair and be easy to remove from the wheelchair to allow transport of the wheelchair without the patient. As such the electronic should be pluggable and the mount shall not require any special tools.

### Provider

The provider must be able to mount, setup, diagnose and calibrate the ASL110.

# Intended Use

The Intended Use of the ASL110 is to provide “joystick demands” to control the wheelchair features.

# User Needs

The User Needs include the patient’s ability to control the Wheelchair features.

# Product Performance requirements/criteria

The ASL110 shall get power from the Wheelchair’s Control System, MK6, Permobil or Curtis powered wheelchair control system.

The ASL110 shall connect to the ASL110 Display Unit which is a PME 290 derivative. Together, these devices shall provide setup and calibration, and diagnostic information.

The ASL110 shall interface to other Control Systems via the generic 9-Pin DSUB connector.

# Project physician feedback

TBD

# Operational Considerations

TBD

# Regulatory Path

TBD

# Preliminary Device Risks

TBD