| Date | Author | Description |
| --- | --- | --- |
| Jan 4, 2020 | G. Chopcinski | Add #15 Veer Adjustment feature |

# Introduction

This document captures the items fixed.

| Item # | Description | Priority | Resolution | Version |
| --- | --- | --- | --- | --- |
| 1 | Improve the Startup Screen | Medium | Changed the “Fusion” per Christine’s recommendation. Left it remain on the screen a little longer; about 1/2 second. | 1.0.1 (DSP) |
| 2 | Improve readability using the colors on the Calibration screen per Joe. | Medium | The colors of the two prompts have been changed to Yellow and muted red, almost orange. | 1.0.1 (DSP) |
| 3 | Implement a Neutral Test when returning from Power On/Off, idle state. The chair drives upon switch pressing and shouldn’t | High | Fixed it, now the system recognizes the Out of Neutral condition and waits for the pads to be released. | 1.0.0 (Head Array) |
| 4 | The display does not always turn on. | High |  |  |
| 5 | The display locks up occasionally. | High |  |  |
| 6 | Add a repeat and hold feature to the arrow buttons in the calibration screen | Medium | Short term solution: Changed the maximum pad values from 100 to 75 | 1.0.0 (Head Array) |
| 7 | Add the number to the joystick demand in Calibration Screen | Low |  |  |
| 8 | Add a “high water” mark/tick to the Calibration Screen | Low |  |  |
| 9 | Diagnostic Screen. Remove the Raw Data. Maybe hide them and show them with a special sequence. | Low |  |  |
| 10 | Diagnostic Screen is showing Orange, active pads even when not activated. | Medium | It was changed to always show the Orange or Green color entirely based upon Pad Setting and not activation. Also added was “OFF” when the pad is to No Direction, “off”. | 1.0.0 (DSP) |
| 11 | Add audible indicator to Modes when changing features | Medium |  |  |
| 12 | Add a “Power up in Idle” feature that requires a Buddy Button activation at power up to wake up the Head Array. | Medium |  |  |
| 13 | A setting that I don’t think we have but might be nice – be able to rest on the pad with no driving and press on the pad to initiate a direction – making the speed “ramp-up” adjustable to be a slow or fast ramp up speed | Low | Additional information is needed. |  |
| 14 | Add a “Reset to Default” feature that sets all programmed settings to a known “factory” default | Low |  |  |
| 15 | Add a “Neutral” or “Veer” adjustment feature that adjusts the “No Drive Demand” signal up or down | Med/HIgh |  |  |

Resolved and Presumed Fixed or decided to not implement.

| Item # | Description | Priority | Resolution | Version |
| --- | --- | --- | --- | --- |