Approvals

| Title | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Gary Chopcinski |  |  |
| Director |  |  |  |
| Project Engineer |  |  |  |

Change History

| Revision | Revision Date | Name | Description of Change |
| --- | --- | --- | --- |
| A | 9/13/16 | Gary Chopcinski | Initial Release |

# Introduction

## Purpose and Scope

This document describes the user requirements specification for the ASL 1701 Display Unit, herein, referred to as ASL-1701 or 1701.

## Definitions

### User

The user is the patient, caregiver, or provider.
Note1: The patient becomes the user when he or she controls the power wheelchair through the driver control.
Note 2: By definition, there is only one user of the power wheelchair at any one time.

#### Patient

The patient is the individual for whom the power wheelchair is prescribed, seated in and controlling the power wheelchair through the driver control input device.

#### Caregiver

The caregiver is anyone who assists patient with transfer into/out of the power wheelchair or assists with the operation of/or positioning the power wheelchair. The caregiver either controls the motion of the power wheelchair manually or using the electronic motion controls.

#### Provider

The provider is the initial user of power wheelchair who receives the power wheelchair from the manufacturer.  The provider configures the power wheelchair through either programming its parameters or mechanically adjusting it for the patient and also provides periodic maintenance for the power wheelchair.

### Power Mobility Device (PMD)

A PMD is a mobility device with direct drive motors to propel the drive wheels, on board batteries, a control module to send electric current to the motors and a control input device (joystick, driver control) to send directional and speed commands to the control module.

### Expandable Electronics

An electronic system capable of accommodating motors for mobility, three or more actuators, accepting alternative driver controls, and operating other devices through the driver control input in addition to the power wheelchair functionality.

### Powered Seating

Actuator controlled seating that changes the body's position in space.

### ISTA

International Safe Transit Association.

### HCPCS

Healthcare Common Procedure Coding System (est. 1978; Centers for Medicare and Medicaid Services)

### CMS

Centers for Medicare and Medicaid Services (formerly HCFA)

### PDAC

Pricing, Data Analysis and Coding Contractor (Noridian Administrative Services, LLC; various locations)

### RESNA

Rehabilitation Engineering and Assistive Technology Society of North America (formerly Rehabilitation Engineering Society of North America)

### ADA

Americans with Disabilities Act of 1990 (US)

### DHF

Design History File

### ASL

Adaptive Switch Laboratories

## Reference Documents

Article for Power Mobility Devices - Policy Article - Effective January 2009 (A47122)
Wheelchair Options/Accessories – Policy Article – Effective January 2010 (A47229)
RESNA 2009 Book 9 - Controlled Standards
RESNA 1998 Book 9a – Controlled Standards
LCD for Power Mobility Devices (L27239)
LCD for Wheelchair Options/Accessories (L27223)

# Product Description

The ASL 1701 Display Unit (ASL 1701) is an additional, add-on device intended to convey PMD operational status, operating mode and auditory feedback.  It is primarily intended to provide visual information and provides an interface between the ASL ATOM modules, the power seating system and the wheel chair’s drive system.

# User Characteristics, Interfaces and Environment

## User Characteristics

### Patient

The patient can be either male or female, vary in age though with a typical range from 4 to 75 years old, (pediatric and adolescent populations through adult to geriatric applications). The majority of patients will have a diagnosis classified as a neurologic condition, myopathy, or congenital skeletal deformity. Neurologic conditions include diagnoses such as Spinal Cord Injury (SCI), Paraplegia and Tetraplegia, Cerebral Palsy (CP), Multiple Sclerosis (MS), Post-Polio Syndrome (PPS), Cerebral Vascular Accident (CVA or Stroke), Traumatic Brain Injury (TBI) and Amyotrophic Lateral Sclerosis (ALS). Myopathies include many muscular dystrophies (MD). Congenital skeletal deformities, (present since birth), may result from one of the above diagnoses, or from orthopedic conditions such as Arthrogryposis or Juvenile Rheumatoid Arthritis. Some of these conditions / diagnoses such as MS, ALS and some muscular dystrophies are progressive and will result in diminishing function over time. Disabilities resulting from these diagnoses can result in a loss of physical functions (i.e. muscle function or control), loss of sensory function (i.e. visual or auditory impairment), or loss of Cognitive skills (problem solving or memory difficulties).

### Caregiver

The Caregiver can be a family member, healthcare professional, or assistant/aide. The caregiver will have the cognitive and physical abilities to assist the patient with participation in one or more mobility related activities of daily living, assist with management of the wheelchair, and as required - assist the patient with transfers into / out the wheelchair and positioning within the wheelchair.

### Provider

The Provider is an individual with the knowledge and skills to be able to match appropriate wheelchair technology to identified patient needs, complete initial set-up of the wheelchair to match patient requirements, perform caregiver and patient training on management of all wheelchair functions, and complete troubleshooting of wheelchair faults / malfunction if required. The Provider is an individual with the knowledge and skills to be able to match appropriate wheelchair technology to identified patient needs, complete initial set-up of the wheelchair to match patient requirements, perform caregiver and patient training on management of all wheelchair functions, and complete troubleshooting of wheelchair faults / malfunction if required.

## User Interfaces

### Patient Interface

The patient requires the ability to view this device and/or listen to voice commands.

### Caregiver Interface

The caregiver could utilize this device if the caregiver can comprehend the displayed information and select features.

### Provider Interface

The Provider requires the ability to install, setup and program the ATOM 1701 suitable for the Patient and the Caregiver.

## User Environment

The user will operate the ASL 1701 in many environments. These include, but are not limited to homes, assisted living facilities, nursing homes, vocational settings, educational settings, health care settings, the outdoors (on firm terrain), ADA compliant environments, tight confined spaces and wide open spaces.

## Markets Environment

Countries where this module will be used include, but are not limited to, the USA, Canada, Australia, New Zealand, Japan and Europe.

# Intended Use

## ASL1701URS-IU-1 Information

The ASL 1701 is intended supply information to the Patient and Caregiver.  This information may include operational status, modal information and is intended to enhance control system electronics.

# User Needs

## Patient

### ASL1701URS-UN-1 View/Identify status

The patient needs to be able to view or identify the wheel chair status.

### ASL1701URS-UN-2 Change modes

The patient requires a method to change between driving mode, power seating mode and communication mode.

## Caregiver

### ASL1701URS-CN-1 View/Identify status

The caregiver needs to be able to identify the wheel chair status.

### ASL1701URS-CN-2 Change modes

The caregiver requires a method to change between driving mode, power seating mode and communication mode.

## Provider

### ASL1701URS-PN-1 Configure

The provider must be able to setup the ATOM-1701 for specific needs of the patient and, secondly, the caregiver.

# Design Requirements

## Functional Requirements

### ASL1701URS-FR-x Display information

The ASL-1701 must have a color graphics display and must be sufficiently sized to be viewed from approximately 24 inches away.

As reference, the iPhone 5S is 4” measured diagonally.

### ASL1701URS-FR-x ATOM interface

The ASL-1701 must connect to the ASL ATOM Module which will be enhanced to communicate bi-directionally.

The ASL-1701 will accept either 3 or 4 quadrant digital commands from an ATOM connected device.

Question: I’m not quite sure what I mean by “3 or 4 quadrant” digital commands.

Question/Concern: I don’t think we are going be able to universally support proportional driver controls.  This may be true because we may not have a method to connect proportional signals to all wheel chair electronics.

### ASL1701URS-FR-x Driving

The ASL-1701 must be able to support the wheel chair in driving mode.  The ATOM-1701 will command the wheel chair via Digital Input capabilities of the wheel chair electronics.

The ASL-1701 will also display ‘driving mode’ information on the screen.

### ASL1701URS-FR-x Seating Control

The ASL-1701 must be able to control power seating.  The ASL-1701 will provide control signals via the Digital Input capabilities of the wheel chair electronics.

### ASL1701URS-FR-x Wireless Interface

The ASL-1701shall support a wireless mode.

Question: Are we including a Wireless Interface in the Display Unit?  I think NOT, however, the Display Unit must be able to indicate that the driver control (ATOM) is in Wireless Operation Mode.

### ASL1701URS-FR-x Programmability

The ASL-1701 must have a method to setup the unit specific for each patient, primarily and caregiver, secondarily.

### ASL1701URS-FR-x Programming Security

The ASL-1701 must have a method to allow the provider to setup/program the ASL-1701 only and not the patient nor caregiver.

If possible, it would be nice to allow the patient and/or caregiver, the capability to change the order of features or the appearance of the information.

### ASL1701URS-FR-x Power On/Off

The ASL-1701module shall include power On/Off with power indication.  The source of power is the nominal 24 Volt system and is always available.

Alternatively, the power source to the ASL-1701 may be switched by the ‘host’ electronic system.

Question: How are we getting power?

### ASL1701URS-FR-x Audible Indicator / Voice Out

The ASL-1701 shall include an audible indicator.

Question: Are we expecting Voice Out capability?

### ASL1701URS-FR-10 Reset/Terminate Feature

The ASL-1701 shall support a Reset Function which acts as a Mode change request when the system is Idle/Ready and acts as a “feature stop” request when a feature is active.

## Physical Requirements

### ASL1701URS-PR-x LiNX Digital Interface

The ASL-1701 requires a LiNX interface connection.  Ultimately, the ASL-1701 must be physically connected to the LiNX Bus using the LiNX Connection scheme.

Question: I don’t think we want to be on the LiNX control bus, but this might be the method to get power.

### ASL1701URS-PR-x Mounting

The ASL-1701 needs to be able to mount the unit to the wheel chair in view of the patient.

### ASL1701URS-PR-x Switch Ports

The ASL-1701 must support two Mono Jack Switch ports to support the Reset/Termination feature and ????.

Question: How many Switch Sockets (Ports) do we want on this device? Mono or Stereo?

## Regulatory Compliance Requirements

### ASL1701URS-RR-x Serial Number Label

The ASL-1701 Display Unit needs to be labeled with a serial number label.

### ASL1701URS-RR-x Model Number

The ASL-1701 Display Unit needs to be labeled with a model number label.

### ASL1701URS-RR-x Cable Label

The ASL-1701 Display Unit cable needs to be labeled indicating type of module.

### ASL1701URS-RR-x Electrostatic Discharge Immunity

The ASL-1701 Display Unit needs to be immune to electrostatic discharge.

## Risk Management Requirements

### ASL1701URS-RM-1 Contact temperature hazard

The unit can be touched by the users, therefore the unit (case) shall remain at an acceptable temperature level per regulations.  The design of the unit does not include heat generating componenents, there for, this should not be a consideration nor potential hazard.

### ASL1701URS-RM-1 Loss of use hazard

The ASL-1701 shall be designed to reduce features loss of use.

vacare Corporation One Invacare Way Elyria, OH 44036; User Requirement Specification.; P/N: ASL104URS; Page 8of 8. Title: ASL 104 Atom Electronic Head Array User Requirement Specification; DHF Number: See Power Wheel Chair DHF.

User Requirement Specification Rev AInvacare Confidential