# ASL110 Head Array driving commands assessment

By: G. Chopcinski and Trevor Parsh						Date: 20-Dec-2019

# Introduction

The ASL110 Head Array is a standard 3-pad driver apparatus designed to control the features of a Power Wheel Chair such as the Invacare TDX-SP2 using LiNX Electronics. The ASL110 features both Digital Demands and Proportional Demands.

# Driving Issue

The ASL110 has a tendency to drive veering to the right when only a forward demand is requested. The advantage of the ASL110 Head Array is that the Left, Center and Right pads can be activated entirely separately meaning that one pad and only may be activated at a time; a joystick does not have that exclusivity demand feature.

This document intends to identify the root-cause of the “bad” driving behavior.

# Investigation Bottom Line

- The Reference Voltage at Pin 3 of the 9-Pin connector, when plugged into the Input Module, is being pulled-down by 113 milli-volts.

- There is a signal being introduced in the opposite demand-pin of 5.13% by the LiNX system.

Conclusion: The LiNX system is causing the veering of the head array. We can minimize the effects, but I don’t think we can totally eliminate them without better control of the Reference Voltage and help by DCL.

Editorial: The LiNX Input System should not be causing items 1 and 2 and those need to be addressed by DCL.

# Investigation Details

Testing was performed on a prototype ASL110 Head Array with Rev 3 board and version 0.2.2 firmware.

## Issue 1: Pin 3 Loading

One of the issues is the inconsistent Pin 3 Reference Voltage. The range of the voltage supplied by the Input Module is 11.6 volts to 12.4 volts, according to the LiNX Input Module documentation; we’ll call this “LiNX Supplied Voltage”. The circuit on the ASL110 is designed to use a voltage divider to reduce the LiNX Supplied Voltage by 50% and supply it to Pin 3 as the Reference Voltage. The ASL110 circuit also uses the LiNX Supplied Voltage as the Vcc and Vref supply to the Digital-To-Analog Converters (DAC’s) through IC1 which drive the Proportional (Joystick) signals at Pins 1 and 2.

We discovered that the voltage across IC1 on the ASL110 Board drops a nominal 1.3 volts across but that voltage varies based upon the LiNX Supplied Voltage. We reduced any negative influence of this device by removing IC1 and jumpering across the device. This ensures that the only difference between the voltage supplied to the DAC’s Vref/Vcc and the voltage supplied to the voltage divider that supplies voltage to pin 3 is PCB trace length.

On the bench, the 12 Volt supplied voltage was varied from 11.6 to 12.4 and the voltage measured at the voltage divider (Pin 3) was 50% +/- 5 milli-volts.

We then connected the Head Array to the LiNX Input Module via the 9-Pin DSub connector. The LiNX system programmed with multiple drive functions using the Input Module as the driver control. The Input Module programmed to use the 9-Pin connector as a Proportional Joystick. The Neutral Window set to 25%.

Per the LiNX Input Module specification, Pin 3 is used as reference and Pins 1 and 2 provide proportional signals. The reference voltage must be 6 volts +/- 1 Volt and Pins 1 and 2 neutral Voltage must have the same value as Pin 3’s Reference Voltage as neutral. Then, +/- 1.2 volts to command full joystick demand on Pins 1 and 2 when out of neutral.

With the Head Array plugged into the LiNX system, the Input Module supplied 12.31 volts and Pin 3 was measured at 6.02 Volts. On the bench, we set the power supply to 12.31 (same as Input Module) and Pin 3 was measured at 6.15 Volts. The difference of the loaded-down reference voltage at Pin 3 is 113 milli-volts. Because the DAC’s are using the 12.31 voltage as Vref, this situation immediately introduces a 10.6% “away from neutral” signal as viewed by the Input Module.

Possible Fix 1:
If we assume that the Input Module will always load Pin 3 at a nominal 110 milli-volts, we can reduce the DAC’s neutral command to match Pin 3. We did this and connected the Head Array to the LiNX system once again. With power applied and with no pads activated, we measured the voltage between Pins 1 and 3 and then Pins 2 and 3. The voltage was < 3 milli-volts and should not cause any drive demand. This means that this Head Array is “calibrated” to the loaded-down Pin 3 Reference Voltage.

Possible Fix 2:

Use much smaller resistance values for the voltage divider network, or put a voltage buffer between the voltage diver network and Pin 3, or use a tightly regulated step-down voltage regulator in place of the voltage divider network and also have a regulator for the DAC that holds a constant voltage. We suggest using a voltage buffer.

## Issue 2. 5.13% opposite channel demand

With the Head Array connected to the LiNX system and driving as expected, there is a signal introduced to the “not used” channel. This is identified using LiNX Diagnostics.

When we activate the Center Pad, there is a 5.13% signal introduced on the Turn Demand. If we change the drive function for the Center Pad to drive in reverse and activate the Center Pad, the turn demand shows 5.13%. Whenever the Center Pad is active, the Turn Demand is 5.13%. The value doesn’t go negative and it does not change even the Center Pad changes proportionally. If the Speed Demand is minimal, i.e. 10%, the Turn Demand is 5.13%. If the Speed Demand is 100%, the Turn Demand is 5.13%. The 5.13% is very consistent.

When we activate the Right Pad, there is a -5.13% signal introduced into Speed demand. When we activate the Left Pad, there is a -5.13% signal introduced into the Speed Demand. It doesn’t change to positive and it is always -5.13% and doesn’t change even though the Turn Demand changes proportionally.

To confirm signals going to the Input Module in this situation, the voltage was measured from Pin 3 (reference) to the opposite channel pin, i.e. Turn or Speed. The voltages were < 3 milli-volts.

Conclusion: The 5.13% opposite channel demand is NOT being initiated by the ASL110 Head Array.

Possible Firmware Patch to Fix Hardware:
We were able to get the veering to go away, but since we do not understand what the IN500 module is doing/requires, we do not know if the change will work for every unit.  What we did was set the neutral voltage on both Pins 1 and 2 to be 80 mV less than the voltage measured at Pin3 of the 9-pin DSub connector. 60 mV under almost gets rid of the veering. 80 mV under gets rid of it entirely.