Approvals

| Title | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Gary Chopcinski |  |  |
| Director |  |  |  |
| Project Engineer |  |  |  |

# Introduction

This document describes the user requirements specification for the ASL110 Proportional Head Array with Display unit.

# Definitions

## User

A user is the patient, caregiver, or provider.

### Patient

The patient is the individual for whom the power wheelchair is prescribed, seated in and controlling the power wheelchair through the driver control input device.

### Caregiver

The caregiver is anyone who assists patient with transfer into/out of the power wheelchair or assists with the operation of/or positioning the power wheelchair. The caregiver either controls the motion of the power wheelchair manually or using the electronic motion controls.

### Provider

The provider is the initial user of power wheelchair who receives the power wheelchair from the manufacturer.  The provider configures the power wheelchair through either programming its parameters or mechanically adjusting it for the patient and also provides periodic maintenance for the power wheelchair.

## Power Mobility Device (PMD), a.k.a. Wheelchair

A PMD is a mobility device with direct drive motors to propel the drive wheels, on board batteries, a control module to send electric current to the motors and a control input device (joystick, driver control) to send directional and speed commands to the control module.

## Expandable Electronics

An electronic system capable of accommodating motors for mobility, three or more actuators, accepting alternative driver controls, and operating other devices through the driver control input in addition to the power wheelchair functionality.

## Powered Seating

Actuator controlled seating that changes the body's position in space.

## ISTA

International Safe Transit Association.

## HCPCS

Healthcare Common Procedure Coding System (est. 1978; Centers for Medicare and Medicaid Services)

## CMS

Centers for Medicare and Medicaid Services (formerly HCFA)

## PDAC

Pricing, Data Analysis and Coding Contractor (Noridian Administrative Services, LLC; various locations)

## RESNA

Rehabilitation Engineering and Assistive Technology Society of North America (formerly Rehabilitation Engineering Society of North America)

## ADA

Americans with Disabilities Act of 1990 (US)

## DHF

Design History File

## ASL

Adaptive Switch Laboratories

# Reference Documents

Article for Power Mobility Devices - Policy Article - Effective January 2009 (A47122)
Wheelchair Options/Accessories – Policy Article – Effective January 2010 (A47229)
RESNA 2009 Book 9 - Controlled Standards
RESNA 1998 Book 9a – Controlled Standards
LCD for Power Mobility Devices (L27239)
LCD for Wheelchair Options/Accessories (L27223)

# Product Description

The ASL110 Head Array with Display is s device mounted to a Wheelchair and provides the user/patient with 3 pads to control the wheelchair features. The Display provides Patient feedback and programming and setup capabilities.

# User Characteristics, Interfaces and Environment

## User Characteristics

### Patient

The patient can be either male or female, vary in age though with a typical range from 4 to 75 years old, (pediatric and adolescent populations through adult to geriatric applications). The majority of patients will have a diagnosis classified as a neurologic condition, myopathy, or congenital skeletal deformity. Neurologic conditions include diagnoses such as Spinal Cord Injury (SCI), Paraplegia and Tetraplegia, Cerebral Palsy (CP), Multiple Sclerosis (MS), Post-Polio Syndrome (PPS), Cerebral Vascular Accident (CVA or Stroke), Traumatic Brain Injury (TBI) and Amyotrophic Lateral Sclerosis (ALS). Myopathies include many muscular dystrophies (MD). Congenital skeletal deformities, (present since birth), may result from one of the above diagnoses, or from orthopedic conditions such as Arthrogryposis or Juvenile Rheumatoid Arthritis. Some of these conditions / diagnoses such as MS, ALS and some muscular dystrophies are progressive and will result in diminishing function over time. Disabilities resulting from these diagnoses can result in a loss of physical functions (i.e. muscle function or control), loss of sensory function (i.e. visual or auditory impairment), or loss of Cognitive skills (problem solving or memory difficulties).

### Caregiver

The Caregiver can be a family member, healthcare professional, or assistant/aide. The caregiver will have the cognitive and physical abilities to assist the patient with participation in one or more mobility related activities of daily living, assist with management of the wheelchair, and as required - assist the patient with transfers into / out the wheelchair and positioning within the wheelchair.

### Provider

The Provider is an individual with the knowledge and skills to be able to match appropriate wheelchair technology to identified patient needs, complete initial set-up of the wheelchair to match patient requirements, perform caregiver and patient training on management of all wheelchair functions, and complete troubleshooting of wheelchair faults / malfunction if required. The Provider is an individual with the knowledge and skills to be able to match appropriate wheelchair technology to identified patient needs, complete initial set-up of the wheelchair to match patient requirements, perform caregiver and patient training on management of all wheelchair functions, and complete troubleshooting of wheelchair faults / malfunction if required.

## User Interfaces

### Patient Interface

The patient requires the ability to view this device and/or listen to voice commands.

### Caregiver Interface

The caregiver could utilize this device if the caregiver can comprehend the displayed information and select features.

### Provider Interface

The Provider requires the ability to install, setup and program the ASL110 Fusion suitable for the Patient and the Caregiver.

## User Environment

The user will operate the ASL110 Fusion in many environments. These include, but are not limited to homes, assisted living facilities, nursing homes, vocational settings, educational settings, health care settings, the outdoors (on firm terrain), ADA compliant environments, tight confined spaces and wide open spaces.

## Markets Environment

Countries where this module will be used include, but are not limited to, the USA, Canada, Australia, New Zealand, Japan and Europe.

# Intended Use

## ASL110-IU-01 Control

The ASL110 Fusion shall control the wheelchair’s features when connected to the wheelchair’s control system.

## ASL110-IU-02 User Feedback

The ASL110 Fusion shall supply visual and audible feedback to the users.  This information may include operational status and modal information.

## ASL110-IU-03 Programmability

The ASl110 Fusion shall be capable of setting up (programming) the operation suitable for the users’ needs.

# User’s Needs

## Patient Needs

### ASL110-PN-01 Patient Wheelchair Control

The patient needs to control the wheelchair’s features using three pads.

### ASL110-PN-02 Patient Proportional Control

The patient needs to control the appropriate wheelchair’s features using proportional control based upon the amount of pressure applied to each pad.

### ASL110-PN-03 Patient Digital Control

The patient needs to control the appropriate wheelchair’s features with a simple Go/No-go schema, i.e. Digital Control.

### ASL110-PN-04 Patient Feature Control

The patient needs to be able to select features and then activate the feature. This is typically performed using a Buddy Switch located near the patient’s head.

### ASL110-PN-05 Patient Feedback

The patient needs to be able to determine the selected feature as well as status information.

### ASL110-PR-06 Bluetooth control

The patient needs to be able to control a finite set of ASL Bluetooth compatible devices.

## Caregiver

### ASL110-CN-06 Caregiver control

The caregiver needs to be able to select the active feature.

## Provider

### ASL110-PN-07 Provider control

The provider needs to be able to program and setup ASL110 Head Array and Display to accommodate physical capabilities of the patient.

# Design Requirements

## ASL110-DR-01

The ASL110 Head Array shall mount to the wheelchair using current mounting methods available on Invacare, and Permobil wheelchairs.

## ASL110-DR-02

The ASL110 Head Array with Display shall be environmentally safe from normal, expected moisture and electrical interference.

## ASL110-DR-03

The ASL110 Head Array shall utilize the (existing) ATOM technology.

## ASL110-DR-04

The ASL110 Head Array shall be connected to the Wheel chair’s electronics via the standard 9-pin DSub connection utilizing the Proportional Speed and Direction “joystick” demands signals of the LiNX Input Module.

## ASL110-DR-05

The ASL110 Head Array shall support the LiNX Next Function and Next Profile utilizing Pin 6 of the 9-Pin DSub connection of the LiNX Input Module.

## ASL110-DR-06

The ASL110 Head Array shall supply the power and communication interface via a dedicated connection between the ASL110 Head Array and the Display unit.

## ASL110-DR-07

The ASL110 Head Array shall provide the Provider the capability to change the operating characteristics of the ASL110 Head Array including Pad Direction, Pad Disabling, Digital vs Proportional operation, diagnostic information and adjustable Proportional settings.

## ASL110-DR-08

The ASL110 Head Array shall support an “Idle” mode that disables the operation of the pads. This allows the pads to be used as a head rest and prevents wheelchair operation. The mode is controllable by the patient.

## ASL110-DR-09

The ASL110 Head Array with Display shall support audible indications of the active feature and activating the feature. The audible indicator can be enabled or disabled.

## ASL110-DR-10

The ASL110 Head Array shall be identified with a Model Number and Serial Number.